## Noise Figure
Frequency [Hz]	Frequency [MHz]	Power Measured, Noise Source Off [dBm]	Power Measured, Noise Source On [dBm]	Y-Factor	Noise Figure [dB]
0	0	-11.99130917	-11.75434494	0.23696423	27.6319943453
217391.3043	0.2173913043	-27.30935097	-27.07202148	0.23732949	27.625120931
434782.6087	0.4347826087	-29.95539856	-29.71936607	0.23603249	27.6495745012
652173.913	0.652173913	-32.98615265	-32.75317383	0.23297882	27.7076685428
869565.2174	0.8695652174	-37.06858444	-36.8364563	0.23212814	27.7239841767
1086956.522	1.086956522	-42.22130585	-41.98287201	0.23843384	27.6044018433
1304347.826	1.304347826	-48.47438812	-48.27235413	0.20203399	28.3421901068
1521739.13	1.52173913	-55.87171555	-55.62693405	0.2447815	27.4870916011
1739130.435	1.739130435	-62.80449295	-62.23963547	0.56485748	23.6929648295
1956521.739	1.956521739	-66.6144104	-65.34380341	1.27060699	19.806949123
2173913.043	2.173913043	-67.75583649	-65.85919189	1.8966446	17.735203416
2391304.348	2.391304348	-69.43760681	-66.32976532	3.10784149	14.9270596565
2608695.652	2.608695652	-70.7430191	-66.56934357	4.17367553	13.0399610915
2826086.957	2.826086957	-70.80683899	-66.82597351	3.98086548	13.3566453619
3043478.261	3.043478261	-71.65588379	-66.65357971	5.00230408	11.7675161394
3260869.565	3.260869565	-72.16804504	-66.57684326	5.59120178	10.9313010193
3478260.87	3.47826087	-72.80612183	-66.22361755	6.58250428	9.6146530657
3695652.174	3.695652174	-73.09596252	-66.27230072	6.8236618	9.307959069
3913043.478	3.913043478	-73.765625	-65.95291138	7.81271362	8.0928854264
4130434.783	4.130434783	-73.80515289	-65.55252075	8.25263214	7.5707927352
4347826.087	4.347826087	-74.30310059	-65.44319153	8.85990906	6.8649966571
4565217.391	4.565217391	-74.50592041	-65.34928131	9.1566391	6.5256111569
4782608.696	4.782608696	-74.59273529	-65.21492004	9.37781525	6.2747319235
5000000	5	-74.65511322	-65.34703064	9.30808258	6.3536449538
5217391.304	5.217391304	-74.70994568	-65.05264282	9.65730286	5.9600739702
5434782.609	5.434782609	-74.88475037	-65.04060364	9.84414673	5.7510955677
5652173.913	5.652173913	-74.99346161	-65.13951111	9.8539505	5.7401595121
5869565.217	5.869565217	-74.9924469	-65.15223694	9.84020996	5.7554878176
6086956.522	6.086956522	-74.87207031	-64.95684814	9.91522217	5.6718754689
6304347.826	6.304347826	-74.95466614	-64.87095642	10.08370972	5.4846629272
6521739.13	6.52173913	-74.82497406	-64.69334412	10.13162994	5.4315626774
6739130.435	6.739130435	-74.38682556	-64.67814636	9.7086792	5.9025047227
6956521.739	6.956521739	-74.11655426	-64.50367737	9.61287689	6.0099216787
7173913.043	7.173913043	-74.16059113	-64.48991394	9.67067719	5.9450795847
7391304.348	7.391304348	-73.55000305	-64.40322113	9.14678192	6.5368325272
7608695.652	7.608695652	-73.69324493	-64.38274384	9.31050109	6.3509052439
7826086.957	7.826086957	-73.96211243	-64.41091919	9.55119324	6.0792372099
8043478.261	8.043478261	-73.85462952	-64.6634903	9.19113922	6.4863640232
8260869.565	8.260869565	-73.31845093	-64.33067322	8.98777771	6.7183371703
8478260.87	8.47826087	-73.68213654	-64.06374359	9.61839295	6.0037290505
8695652.174	8.695652174	-73.45938873	-64.29856873	9.16082	6.5208527024
8913043.478	8.913043478	-72.86204529	-64.22408295	8.63796234	7.1211098538
9130434.783	9.130434783	-72.8664093	-64.34143829	8.52497101	7.2522867523
9347826.087	9.347826087	-73.16951752	-64.07247925	9.09703827	6.5935145452
9565217.391	9.565217391	-72.32609558	-63.77681732	8.54927826	7.2240208099
9782608.696	9.782608696	-72.36161804	-63.98671341	8.37490463	7.4273731381
10000000	10	-72.72357941	-63.92882538	8.79475403	6.9399736988
10217391.3	10.2173913	-72.76444244	-64.15724182	8.60720062	7.1567682954
10434782.61	10.43478261	-72.85693359	-63.77352142	9.08341217	6.6090571077
10652173.91	10.65217391	-73.01094055	-63.99614334	9.01479721	6.6874275214
10869565.22	10.86956522	-72.47853088	-64.01753235	8.46099853	7.3268019224
11086956.52	11.08695652	-72.65557098	-63.99403	8.66154098	7.0938050237
11304347.83	11.30434783	-72.9265976	-63.84386063	9.08273697	6.6098274507
11521739.13	11.52173913	-73.00440979	-64.33384705	8.67056274	7.0833637166
11739130.43	11.73913043	-73.00048828	-63.91174316	9.08874512	6.6029732803
11956521.74	11.95652174	-73.14477539	-64.2748642	8.86991119	6.8535016725
12173913.04	12.17391304	-73.51478577	-64.18322754	9.33155823	6.3270600698
12391304.35	12.39130435	-73.84142303	-64.15673065	9.68469238	5.9293727008
12608695.65	12.60869565	-73.99526978	-64.03100586	9.96426392	5.6172998738
12826086.96	12.82608696	-74.09633636	-63.99313354	10.10320282	5.4630550376
13043478.26	13.04347826	-74.36126709	-63.91299438	10.44827271	5.0822344478
13260869.57	13.26086957	-74.02961731	-64.03848267	9.99113464	5.5874264236
13478260.87	13.47826087	-74.24146271	-63.80094528	10.44051743	5.0907592962
13695652.17	13.69565217	-74.10066986	-63.78173065	10.31893921	5.2246016779
13913043.48	13.91304348	-74.26777649	-63.60549164	10.66228485	4.8475698304
14130434.78	14.13043478	-74.00718689	-63.75497437	10.25221252	5.2982219506
14347826.09	14.34782609	-73.93209839	-63.56461334	10.36748505	5.1711134137
14565217.39	14.56521739	-74.26480103	-63.82841873	10.4363823	5.0953053795
14782608.7	14.7826087	-74.39178467	-63.871521	10.52026367	5.0031713099
15000000	15	-74.59461975	-63.93060303	10.66401672	4.8456753391
15217391.3	15.2173913	-74.71038818	-63.99493408	10.7154541	4.7894402144
15434782.61	15.43478261	-74.63832855	-63.94477081	10.69355774	4.8133713333
15652173.91	15.65217391	-74.81771088	-63.96785736	10.84985352	4.6427938911
15869565.22	15.86956522	-74.81835938	-64.06208038	10.756279	4.7448513865
16086956.52	16.08695652	-75.04400635	-64.16493225	10.8790741	4.6109648908
16304347.83	16.30434783	-75.07546234	-64.09477234	10.98069	4.5004253719
16521739.13	16.52173913	-75.33847046	-64.23958588	11.09888458	4.3721327128
16739130.43	16.73913043	-75.23096466	-64.14215088	11.08881378	4.3830523479
16956521.74	16.95652174	-75.4506073	-64.05026245	11.40034485	4.0462249358
17173913.04	17.17391304	-75.24313354	-64.34202576	10.90110778	4.5869769407
17391304.35	17.39130435	-75.53799438	-64.47026825	11.06772613	4.4059243372
17608695.65	17.60869565	-75.43787384	-64.45251465	10.98535919	4.4953515777
17826086.96	17.82608696	-75.32865143	-64.11492157	11.21372986	4.2477568338
18043478.26	18.04347826	-75.39024353	-63.9429512	11.44729233	3.9956323742
18260869.57	18.26086957	-75.46311951	-64.25050354	11.21261597	4.2489618574
18478260.87	18.47826087	-75.26559448	-64.29623413	10.96936035	4.5127387495
18695652.17	18.69565217	-75.30545044	-63.94272614	11.3627243	4.0867972224
18913043.48	18.91304348	-75.34015656	-64.38813782	10.95201874	4.5315915047
19130434.78	19.13043478	-75.08746338	-64.49475098	10.5927124	4.9237341271
19347826.09	19.34782609	-75.22006989	-64.14842224	11.07164765	4.4016702846
19565217.39	19.56521739	-75.03025818	-64.28623962	10.74401856	4.7582381436
19782608.7	19.7826087	-74.82463074	-64.30835724	10.5162735	5.007550102
20000000	20	-74.48833466	-64.43331909	10.05501557	5.5164892936
20217391.3	20.2173913	-74.41654205	-63.98868561	10.42785644	5.1046799069
20434782.61	20.43478261	-74.49336243	-64.25804138	10.23532105	5.3168771122
20652173.91	20.65217391	-73.98191833	-63.91373825	10.06818008	5.5018849127
20869565.22	20.86956522	-73.74051666	-64.16559601	9.57492065	6.0525596729
21086956.52	21.08695652	-73.82462311	-64.30021667	9.52440644	6.1093763818
21304347.83	21.30434783	-73.49871063	-64.20330048	9.29541015	6.3680037134
21521739.13	21.52173913	-73.64199066	-64.33314514	9.30884552	6.3527806666
21739130.43	21.73913043	-73.95728302	-64.20158386	9.75569916	5.8498884238
21956521.74	21.95652174	-74.12381744	-64.17578888	9.94802856	5.6353595291
22173913.04	22.17391304	-74.09098816	-64.28997803	9.80101013	5.7992482524
22391304.35	22.39130435	-74.33403778	-64.19422913	10.13980865	5.4225061821
22608695.65	22.60869565	-74.46742249	-64.30345917	10.16396332	5.3957697545
22826086.96	22.82608696	-74.70526886	-64.22190857	10.48336029	5.0436838837
23043478.26	23.04347826	-74.82462311	-64.01423645	10.81038666	4.6858143178
23260869.57	23.26086957	-74.91353607	-64.27472687	10.6388092	4.8732568406
23478260.87	23.47826087	-74.80521393	-64.27880859	10.52640534	4.9964322379
23695652.17	23.69565217	-74.94657898	-64.27531433	10.67126465	4.8377476027
23913043.48	23.91304348	-75.04017639	-64.33308411	10.70709228	4.7985777486
24130434.78	24.13043478	-74.82086945	-64.32089996	10.49996949	5.0254461172
24347826.09	24.34782609	-74.75188446	-64.32442474	10.42745972	5.1051161612
24565217.39	24.56521739	-75.0558548	-64.39015961	10.66569519	4.8438393296
24782608.7	24.7826087	-74.99482727	-64.07180023	10.92302704	4.5631242375
25000000	25	-74.92945862	-64.22285461	10.70660401	4.7991113652
25217391.3	25.2173913	-74.98973083	-64.14735413	10.8423767	4.6509412188
25434782.61	25.43478261	-74.83776093	-64.14304352	10.69471741	4.8121036174
25652173.91	25.65217391	-74.78643799	-64.27047729	10.5159607	5.0078933837
25869565.22	25.86956522	-74.85609436	-64.14464569	10.71144867	4.7938170188
26086956.52	26.08695652	-75.16249084	-64.2201004	10.94239044	4.5420616502
26304347.83	26.30434783	-74.81100464	-64.3344574	10.47654724	5.0511669481
26521739.13	26.52173913	-74.73356628	-64.1029129	10.63065338	4.8821839897
26739130.43	26.73913043	-74.84461975	-64.18247986	10.66213989	4.8477284053
26956521.74	26.95652174	-74.98605347	-64.21537781	10.77067566	4.7291365871
27173913.04	27.17391304	-75.10969543	-64.25626373	10.8534317	4.6388952712
27391304.35	27.39130435	-74.99898529	-64.12654114	10.87244415	4.6181849981
27608695.65	27.60869565	-74.88171387	-63.9965744	10.88513947	4.6043604775
27826086.96	27.82608696	-74.7901535	-64.0511322	10.7390213	4.7636954786
28043478.26	28.04347826	-74.78247833	-64.19047546	10.59200287	4.924511481
28260869.57	28.26086957	-75.02671814	-64.0875473	10.93917084	4.5455632073
28478260.87	28.47826087	-75.03601837	-64.24711609	10.78890228	4.7092479875
28695652.17	28.69565217	-74.9397583	-64.21068573	10.72907257	4.7745618661
28913043.48	28.91304348	-75.15668488	-64.00536346	11.15132142	4.3153104518
29130434.78	29.13043478	-74.87994385	-64.23109436	10.64884949	4.862269174
29347826.09	29.34782609	-74.73290253	-64.08595276	10.64694977	4.864347963
29565217.39	29.56521739	-74.80058289	-64.02570343	10.77487946	4.7245487914
29782608.7	29.7826087	-74.79386139	-64.13220978	10.66165161	4.8482625489
30000000	30	-74.64498138	-64.35528564	10.28969574	5.256851933
30217391.3	30.2173913	-74.77867126	-64.28220367	10.49646759	5.0292908171
30434782.61	30.43478261	-74.67179871	-64.27862549	10.39317322	5.1428344598
30652173.91	30.65217391	-74.77639008	-63.9835434	10.79284668	4.7049449211
30869565.22	30.86956522	-74.89933777	-63.98084641	10.91849136	4.5680591277
31086956.52	31.08695652	-74.76811981	-64.20365906	10.56446075	4.954695757
31304347.83	31.30434783	-74.87311554	-64.05903625	10.81407929	4.6817877113
31521739.13	31.52173913	-74.64003754	-64.11010742	10.52993012	4.9925650199
31739130.43	31.73913043	-74.93582916	-63.99311066	10.9427185	4.541704873
31956521.74	31.95652174	-74.61646271	-63.92526245	10.69120026	4.8159485557
32173913.04	32.17391304	-74.94195557	-63.90156937	11.0403862	4.4355916826
32391304.35	32.39130435	-75.10417938	-64.2824707	10.82170868	4.6734692682
32608695.65	32.60869565	-74.88896942	-64.289711	10.59925842	4.9165629456
32826086.96	32.82608696	-75.05574799	-64.25352478	10.80222321	4.6947172035
33043478.26	33.04347826	-74.76361847	-64.16029358	10.60332489	4.9121086369
33260869.57	33.26086957	-74.81978607	-64.02172089	10.79806518	4.699252446
33478260.87	33.47826087	-74.95257568	-64.1255722	10.82700348	4.6676970402
33695652.17	33.69565217	-74.99253845	-64.19616699	10.79637146	4.7010999316
33913043.48	33.91304348	-74.99848938	-64.25237274	10.74611664	4.7559470751
34130434.78	34.13043478	-75.04120636	-64.05670166	10.9845047	4.4962800774
34347826.09	34.34782609	-74.85307312	-64.30054474	10.55252838	4.9677785315
34565217.39	34.56521739	-74.91560364	-64.1411972	10.77440644	4.7250649991
34782608.7	34.7826087	-74.96870422	-64.14465332	10.8240509	4.6709157737
35000000	35	-74.72984314	-64.21887207	10.51097107	5.0133695682
35217391.3	35.2173913	-74.93263245	-64.24980164	10.68283081	4.8250991826
35434782.61	35.43478261	-75.00539398	-64.01139069	10.99400329	4.4859596778
35652173.91	35.65217391	-75.22889709	-64.30021667	10.92868042	4.556973907
35869565.22	35.86956522	-75.22747803	-64.3782959	10.84918213	4.6435254388
36086956.52	36.08695652	-75.31874847	-64.2589798	11.05976867	4.4145575667
36304347.83	36.30434783	-74.960495	-64.29234314	10.66815186	4.8411521992
36521739.13	36.52173913	-74.94078064	-64.23854828	10.70223236	4.8038892671
36739130.43	36.73913043	-75.1881485	-64.26138306	10.92676544	4.5590571412
36956521.74	36.95652174	-75.41871643	-64.30053711	11.11817932	4.3512176304
37173913.04	37.17391304	-75.26013947	-64.50299072	10.75714875	4.7439018678
37391304.35	37.39130435	-75.44788361	-64.22431946	11.22356415	4.2371190419
37608695.65	37.60869565	-75.53401184	-64.47686005	11.05715179	4.4173969731
37826086.96	37.82608696	-75.39095306	-64.49191284	10.89904022	4.5892274237
38043478.26	38.04347826	-75.50746155	-64.28083801	11.22662354	4.2338100883
38260869.57	38.26086957	-75.68080139	-64.60081482	11.07998657	4.3926253322
38478260.87	38.47826087	-75.31052399	-64.30671692	11.00380707	4.4753097288
38695652.17	38.69565217	-75.63084412	-64.45336914	11.17747498	4.2869909512
38913043.48	38.91304348	-75.48822784	-64.48513794	11.0030899	4.4760887277
39130434.78	39.13043478	-75.61436462	-64.68629456	10.92807006	4.5576378858
39347826.09	39.34782609	-75.5139389	-64.69780731	10.81613159	4.6795499268
39565217.39	39.56521739	-75.74758148	-64.70810699	11.03947449	4.4365812797
39782608.7	39.7826087	-75.91065979	-64.66529846	11.24536133	4.2135479416
40000000	40	-75.89892578	-64.61640167	11.28252411	4.173382848
40217391.3	40.2173913	-75.5331192	-64.79053497	10.74258423	4.7598044667
40434782.61	40.43478261	-76.05297089	-64.86800385	11.18496704	4.2788810244
40652173.91	40.65217391	-75.79703522	-64.89900208	10.89803314	4.5903236373
40869565.22	40.86956522	-75.85343933	-64.85600281	10.99743652	4.4822298879
41086956.52	41.08695652	-76.15264893	-64.71273804	11.43991089	4.0035841124
41304347.83	41.30434783	-76.09310913	-64.66358185	11.42952728	4.0147717413
41521739.13	41.52173913	-75.92241669	-64.79222107	11.13019562	4.3381961556
41739130.43	41.73913043	-76.0559082	-65.08765411	10.96825409	4.513941213
41956521.74	41.95652174	-76.38733673	-64.71489716	11.67243957	3.7535816712
42173913.04	42.17391304	-76.08969116	-65.09364319	10.99604797	4.4837383492
42391304.35	42.39130435	-76.03078461	-65.00233459	11.02845002	4.4485489726
42608695.65	42.60869565	-76.19051361	-65.26438904	10.92612457	4.5597543374
42826086.96	42.82608696	-76.12742615	-65.2719574	10.85546875	4.6366759239
43043478.26	43.04347826	-76.38376617	-65.20870972	11.17505645	4.2896091828
43260869.57	43.26086957	-76.3943634	-65.08128357	11.31307983	4.1403792965
43478260.87	43.47826087	-76.40690613	-65.17927551	11.22763062	4.2327208993
43695652.17	43.69565217	-76.55989838	-65.3822403	11.17765808	4.2867927373
43913043.48	43.91304348	-76.55091858	-65.37236023	11.17855835	4.2858181648
44130434.78	44.13043478	-76.84955597	-65.29286957	11.5566864	3.8779075908
44347826.09	44.34782609	-76.72579193	-65.30142212	11.42436981	4.0203293327
44565217.39	44.56521739	-76.90139771	-65.29016113	11.61123658	3.8192868078
44782608.7	44.7826087	-76.71944427	-65.4553299	11.26411437	4.1932764071
45000000	45	-76.96226501	-65.62431335	11.33795166	4.1135286582
45217391.3	45.2173913	-77.04482269	-65.72602844	11.31879425	4.1342091528
45434782.61	45.43478261	-76.99053192	-65.75630188	11.23423004	4.2255839266
45652173.91	45.65217391	-77.00837708	-65.73703003	11.27134705	4.1854599647
45869565.22	45.86956522	-77.12564087	-65.93618011	11.18946076	4.2740172624
46086956.52	46.08695652	-77.16051483	-65.83976746	11.32074737	4.1321004204
46304347.83	46.30434783	-77.29103088	-65.89839172	11.39263916	4.0545329766
46521739.13	46.52173913	-77.21916962	-66.08417511	11.13499451	4.3329966752
46739130.43	46.73913043	-77.16470337	-66.13128662	11.03341675	4.4431569813
46956521.74	46.95652174	-77.18651581	-66.426651	10.75986481	4.7409368182
47173913.04	47.17391304	-77.39506531	-66.18965912	11.20540619	4.2567621215
47391304.35	47.39130435	-77.29965973	-66.49287415	10.80678558	4.6897413925
47608695.65	47.60869565	-77.6743927	-66.49877167	11.17562103	4.2889979738
47826086.96	47.82608696	-77.63111115	-66.53927612	11.09183503	4.3797762224
48043478.26	48.04347826	-77.66578674	-66.50498962	11.16079712	4.3050483775
48260869.57	48.26086957	-77.73119354	-66.65475464	11.0764389	4.3964731935
48478260.87	48.47826087	-77.9087677	-66.62548065	11.28328705	4.1725585619
48695652.17	48.69565217	-78.04768372	-66.76216125	11.28552247	4.1701434641
48913043.48	48.91304348	-78.27500153	-66.9750824	11.29991913	4.1545920255
49130434.78	49.13043478	-78.09738922	-67.04558563	11.05180359	4.4232004101
49347826.09	49.34782609	-78.24651337	-66.99937439	11.24713898	4.2116260452
49565217.39	49.56521739	-78.36375427	-67.27207947	11.0916748	4.3799499647
49782608.7	49.7826087	-78.37054443	-67.18508911	11.18545532	4.2783525162
50000000	50	-78.46876526	-67.44123077	11.02753449	4.4495429498
50217391.3	50.2173913	-78.53226471	-67.49909973	11.03316498	4.4434302956
50434782.61	50.43478261	-78.48509979	-67.48686981	10.99822998	4.4813679257
50652173.91	50.65217391	-78.25428772	-67.43919373	10.81509399	4.6806812919
50869565.22	50.86956522	-77.93986511	-67.61476135	10.32510376	5.2178061388
51086956.52	51.08695652	-77.88304901	-67.7460556	10.13699341	5.425623363
51304347.83	51.30434783	-78.05119324	-67.72927856	10.32191468	5.2213215224
51521739.13	51.52173913	-78.25597382	-67.96530914	10.29066468	5.2557830094
51739130.43	51.73913043	-78.64247894	-67.93221283	10.71026611	4.7951092947
51956521.74	51.95652174	-78.80003357	-68.29960632	10.50042725	5.0249435697
52173913.04	52.17391304	-79.07688141	-68.22937012	10.84751129	4.6453460328
52391304.35	52.39130435	-79.35343933	-68.47247314	10.88096619	4.6089045584
52608695.65	52.60869565	-79.61909485	-68.52690125	11.0921936	4.3793874158
52826086.96	52.82608696	-79.40223694	-68.73858643	10.66365051	4.84607593
53043478.26	53.04347826	-79.98161316	-68.8060379	11.17557526	4.2890475238
53260869.57	53.26086957	-79.75785828	-69.17069244	10.58716584	4.9298112062
53478260.87	53.47826087	-79.89911652	-69.16566467	10.73345185	4.7697783642
53695652.17	53.69565217	-80.10810089	-69.41433716	10.69376373	4.813146149
53913043.48	53.91304348	-80.20404816	-69.51672363	10.68732453	4.820185829
54130434.78	54.13043478	-80.21821594	-69.360672	10.85754394	4.6344151192
54347826.09	54.34782609	-79.78277588	-69.59672546	10.18605042	5.3713357553
54565217.39	54.56521739	-78.76880646	-69.81038666	8.9584198	6.7519534217
54782608.7	54.7826087	-77.50254822	-69.76555634	7.73699188	8.1837797355
55000000	55	-76.25164032	-69.70982361	6.54181671	9.6668629313
55217391.3	55.2173913	-75.51876831	-69.65032959	5.86843872	10.5528772644
55434782.61	55.43478261	-75.90893555	-69.76594543	6.14299012	10.1863551658
55652173.91	55.65217391	-76.91920471	-70.28158569	6.63761902	9.5441493197
55869565.22	55.86956522	-78.35783386	-70.53009796	7.8277359	8.0748906259
56086956.52	56.08695652	-79.51385498	-70.65022278	8.8636322	6.8607173631
56304347.83	56.30434783	-80.75156403	-70.83927155	9.91229248	5.6751379505
56521739.13	56.52173913	-80.9524765	-70.92736053	10.02511597	5.5496772172
56739130.43	56.73913043	-81.31533051	-70.98511505	10.33021546	5.2121719653
56956521.74	56.95652174	-81.35207367	-71.0994873	10.25258637	5.2978091513
57173913.04	57.17391304	-81.44624329	-71.39131165	10.05493164	5.516582419
57391304.35	57.39130435	-81.48139954	-71.72305298	9.75834656	5.8469279402
57608695.65	57.60869565	-81.73240662	-71.74886322	9.9835434	5.595863872
57826086.96	57.82608696	-82.06419373	-71.8001709	10.26402283	5.2851829807
58043478.26	58.04347826	-82.22335052	-71.93416595	10.28918457	5.2574158598
58260869.57	58.26086957	-82.33327484	-71.94932556	10.38394928	5.152986741
58478260.87	58.47826087	-82.20336151	-72.35509491	9.8482666	5.7464995263
58695652.17	58.69565217	-81.82099915	-72.31001282	9.51098633	6.1244847906
58913043.48	58.91304348	-81.16157532	-72.27873993	8.88283539	6.8386543293
59130434.78	59.13043478	-80.72842407	-72.37924194	8.34918213	7.4574865708
59347826.09	59.34782609	-80.55490875	-72.44475555	8.1101532	7.7388133771
59565217.39	59.56521739	-80.4171524	-72.61538696	7.80176544	8.1060077424
59782608.7	59.7826087	-80.53314972	-72.74004364	7.79310608	8.1163913745
60000000	60	-80.53755188	-72.79734039	7.74021149	8.1799085156
60217391.3	60.2173913	-80.38539886	-72.7652359	7.62016296	8.3246514232
60434782.61	60.43478261	-80.30888367	-73.15041351	7.15847016	8.889442664
60652173.91	60.65217391	-80.29236603	-73.41719818	6.87516785	9.2430435756
60869565.22	60.86956522	-81.29347229	-73.45783234	7.83563995	8.0654275058
61086956.52	61.08695652	-81.4043045	-73.65700531	7.74729919	8.1713883961
61304347.83	61.30434783	-81.87042236	-73.88663483	7.98378753	7.8886763698
61521739.13	61.52173913	-82.00483704	-74.03147888	7.97335816	7.9010815539
61739130.43	61.73913043	-82.48990631	-73.9972229	8.49268341	7.289872752
61956521.74	61.95652174	-82.66181946	-74.33786774	8.32395172	7.4870536062
62173913.04	62.17391304	-83.04972076	-74.39678192	8.65293884	7.1037638674
62391304.35	62.39130435	-83.32966614	-74.47090912	8.85875702	6.8663208972
62608695.65	62.60869565	-83.22769928	-74.67082214	8.55687714	7.2151896478
62826086.96	62.82608696	-83.47102356	-74.96078491	8.51023865	7.2694310074
63043478.26	63.04347826	-83.52866364	-75.02131653	8.50734711	7.2727970522
63260869.57	63.26086957	-83.23939514	-74.92097473	8.31842041	7.4935395959
63478260.87	63.47826087	-83.32798767	-75.17274475	8.15524292	7.6855334145
63695652.17	63.69565217	-83.76759338	-75.24217224	8.52542114	7.2517630795
63913043.48	63.91304348	-84.38884735	-75.6192627	8.76958465	6.9689829879
64130434.78	64.13043478	-84.53482819	-75.75032806	8.78450013	6.9517888662
64347826.09	64.34782609	-84.77827454	-75.97353363	8.80474091	6.9284702849
64565217.39	64.56521739	-85.33409119	-76.13264465	9.20144654	6.4746467769
64782608.7	64.7826087	-85.26881409	-76.33377075	8.93504334	6.7787442797
65000000	65	-85.46437073	-76.58168793	8.8826828	6.8388295866
65217391.3	65.2173913	-85.68286896	-76.76039124	8.92247772	6.7931539764
65434782.61	65.43478261	-85.75898743	-76.70683289	9.05215454	6.6447370221
65652173.91	65.65217391	-85.7155838	-76.93788147	8.77770233	6.9596240491
65869565.22	65.86956522	-86.09086609	-77.22393799	8.8669281	6.8569295862
66086956.52	66.08695652	-85.88470459	-77.09860229	8.7861023	6.9499424684
66304347.83	66.30434783	-86.05462646	-77.01970673	9.03491973	6.6644257777
66521739.13	66.52173913	-86.37687683	-77.61719513	8.7596817	6.9804038059
66739130.43	66.73913043	-86.10379028	-77.59446716	8.50932312	7.2704967371
66956521.74	66.95652174	-86.21642303	-77.73770142	8.47872161	7.3061399304
67173913.04	67.17391304	-86.25178528	-77.74567413	8.50611115	7.2742359468
67391304.35	67.39130435	-86.42797852	-77.90193176	8.52604676	7.2510352596
67608695.65	67.60869565	-86.29515076	-77.96784973	8.32730103	7.4831269127
67826086.96	67.82608696	-86.53401184	-78.22543335	8.30857849	7.5050837179
68043478.26	68.04347826	-86.79637909	-78.24913025	8.54724884	7.2263797577
68260869.57	68.26086957	-86.81448364	-78.47307587	8.34140777	7.4665940492
68478260.87	68.47826087	-86.78945923	-78.5723114	8.21714783	7.6125461349
68695652.17	68.69565217	-86.81301117	-78.55595398	8.25705719	7.5655901191
68913043.48	68.91304348	-86.81610107	-78.93830872	7.87779235	8.015017651
69130434.78	69.13043478	-86.93493652	-78.58622742	8.3487091	7.4580406344
69347826.09	69.34782609	-87.29421234	-78.93768311	8.35652923	7.4488821838
69565217.39	69.56521739	-87.02078247	-79.17084503	7.84993744	8.0483184562
69782608.7	69.7826087	-87.40093231	-79.07899475	8.32193756	7.4894152363
70000000	70	-87.3057785	-79.24319458	8.06258392	7.7951330186
70217391.3	70.2173913	-87.40101624	-79.54468536	7.85633088	8.0406713436
70434782.61	70.43478261	-87.14692688	-79.57946777	7.56745911	8.38845981
70652173.91	70.65217391	-87.52178955	-79.64676666	7.87502289	8.018326699
70869565.22	70.86956522	-87.32141113	-79.6175766	7.70383453	8.2236816497
71086956.52	71.08695652	-87.73530579	-79.98830414	7.74700165	8.1717460122
71304347.83	71.30434783	-87.77687073	-79.9885788	7.78829193	8.1221659143
71521739.13	71.52173913	-87.80528259	-80.00865936	7.79662323	8.1121733841
71739130.43	71.73913043	-87.53786469	-80.2374649	7.30039979	8.7143783833
71956521.74	71.95652174	-87.83144379	-80.26257324	7.56887055	8.3867488521
72173913.04	72.17391304	-87.77987671	-80.36961365	7.41026306	8.5797653699
72391304.35	72.39130435	-87.79582214	-80.32997131	7.46585083	8.5119436268
72608695.65	72.60869565	-87.8584671	-80.63053131	7.22793579	8.8035928007
72826086.96	72.82608696	-87.76051331	-80.76079559	6.99971772	9.0868815973
73043478.26	73.04347826	-87.82221222	-80.74305725	7.07915497	8.987866394
73260869.57	73.26086957	-87.79210663	-80.89209747	6.90000916	9.2118061319
73478260.87	73.47826087	-87.98830414	-80.97000885	7.01829529	9.0636853506
73695652.17	73.69565217	-87.79811096	-81.21259308	6.58551788	9.6107915399
73913043.48	73.91304348	-87.9997406	-81.31587219	6.68386841	9.4851773775
74130434.78	74.13043478	-87.80831909	-81.2771759	6.53114319	9.6805821015
74347826.09	74.34782609	-88.1288681	-81.75141907	6.37744903	9.8792198961
74565217.39	74.56521739	-88.09996033	-81.3502655	6.74969483	9.4015378434
74782608.7	74.7826087	-88.20426178	-81.71173096	6.49253082	9.7302931568
75000000	75	-88.0197525	-81.5684967	6.4512558	9.7835737287
75217391.3	75.2173913	-88.06745148	-81.76560974	6.30184174	9.9777050198
75434782.61	75.43478261	-88.02709198	-81.75540161	6.27169037	10.0171257419
75652173.91	75.65217391	-88.15932465	-81.77290344	6.38642121	9.8675670829
75869565.22	75.86956522	-88.36534119	-82.02555084	6.33979035	9.928208781
76086956.52	76.08695652	-88.25410461	-82.11621094	6.13789367	10.1930893372
76304347.83	76.30434783	-88.22538757	-82.2858963	5.93949127	10.4572724575
76521739.13	76.52173913	-88.06748962	-82.28256989	5.78491973	10.6659581896
76739130.43	76.73913043	-87.91571808	-82.46015167	5.45556641	11.1197738174
76956521.74	76.95652174	-88.01524353	-82.36006165	5.65518188	10.8431789729
77173913.04	77.17391304	-88.03434753	-82.35853577	5.67581176	10.8148684173
77391304.35	77.39130435	-88.14778137	-82.37187195	5.77590942	10.678204048
77608695.65	77.60869565	-87.93049622	-82.59468842	5.3358078	11.2881411252
77826086.96	77.82608696	-88.08961487	-82.86076355	5.22885132	11.4401409085
78043478.26	78.04347826	-88.01345825	-82.8436203	5.16983795	11.5246927449
78260869.57	78.26086957	-88.13181305	-82.78552246	5.34629059	11.2733277119
78478260.87	78.47826087	-87.87252045	-82.75428772	5.11823273	11.599042783
78695652.17	78.69565217	-88.30793762	-83.11578369	5.19215393	11.4926608427
78913043.48	78.91304348	-88.07182312	-83.14775848	4.92406464	11.882393075
79130434.78	79.13043478	-88.00430298	-83.38368988	4.6206131	12.33760396
79347826.09	79.34782609	-88.29016113	-83.07106781	5.21909332	11.4540875415
79565217.39	79.56521739	-88.14643097	-83.28617859	4.86025238	11.9768156338
79782608.7	79.7826087	-88.03510284	-83.64198303	4.39311981	12.6899450049
80000000	80	-88.1505127	-83.29314423	4.85736847	11.9810987362
80217391.3	80.2173913	-88.13879395	-83.30271912	4.83607483	12.0127663105
80434782.61	80.43478261	-88.12400055	-83.42510986	4.69889069	12.2186395347
80652173.91	80.65217391	-88.19859314	-83.8393631	4.35923004	12.7433212223
80869565.22	80.86956522	-88.30750275	-83.78521729	4.52228546	12.488653994
81086956.52	81.08695652	-88.04803467	-83.99880981	4.04922486	13.2433287704
81304347.83	81.30434783	-87.96338654	-83.90058899	4.06279755	13.2209682659
81521739.13	81.52173913	-88.15593719	-83.80482483	4.35111236	12.756142229
81739130.43	81.73913043	-88.09835815	-83.76229095	4.3360672	12.7799412799
81956521.74	81.95652174	-88.21401215	-84.13916016	4.07485199	13.2011468335
82173913.04	82.17391304	-88.21116638	-84.17308807	4.03807831	13.2617262433
82391304.35	82.39130435	-87.99671936	-84.0568924	3.93982696	13.4252443802
82608695.65	82.60869565	-88.26070404	-84.07229614	4.1884079	13.0161279712
82826086.96	82.82608696	-88.30532837	-84.30483246	4.00049591	13.3239844191
83043478.26	83.04347826	-88.44459534	-84.3225174	4.12207794	13.1238302848
83260869.57	83.26086957	-88.25532532	-84.66542053	3.58990479	14.0291146751
83478260.87	83.47826087	-88.33422852	-84.4048233	3.92940522	13.4427348153
83695652.17	83.69565217	-88.28447723	-84.29695892	3.98751831	13.3455654558
83913043.48	83.91304348	-88.26490784	-84.40657043	3.85833741	13.5627749778
84130434.78	84.13043478	-88.4101181	-84.60652924	3.80358886	13.6561889055
84347826.09	84.34782609	-88.27347565	-84.48709869	3.78637696	13.6857297212
84565217.39	84.56521739	-88.35145569	-84.77461243	3.57684326	14.0523637149
84782608.7	84.7826087	-88.39717865	-84.67166138	3.72551727	13.7908646721
85000000	85	-88.42298126	-84.86779022	3.55519104	14.0910246702
85217391.3	85.2173913	-88.23323059	-84.98560333	3.24762726	14.6576452851
85434782.61	85.43478261	-88.50436401	-84.9559021	3.54846191	14.103070764
85652173.91	85.65217391	-88.44972992	-85.35970306	3.09002686	14.9619834583
85869565.22	85.86956522	-88.56999207	-85.04576111	3.52423096	14.146570427
86086956.52	86.08695652	-88.5273056	-85.28408051	3.24322509	14.6660088626
86304347.83	86.30434783	-88.59072113	-85.23058319	3.36013794	14.4464158654
86521739.13	86.52173913	-88.66267395	-85.26670837	3.39596558	14.3801410933
86739130.43	86.73913043	-88.5816803	-85.51759338	3.06408692	15.0130840787
86956521.74	86.95652174	-88.67098999	-85.36219788	3.30879211	14.5422184266
87173913.04	87.17391304	-88.61050415	-85.52159882	3.08890533	14.9641867196
87391304.35	87.39130435	-88.67906189	-85.57054138	3.10852051	14.9257312218
87608695.65	87.60869565	-88.5460434	-85.63848877	2.90755463	15.3279804226
87826086.96	87.82608696	-88.72386169	-85.69380951	3.03005218	15.0805849821
88043478.26	88.04347826	-88.48423004	-85.50791931	2.97631073	15.1882465631
88260869.57	88.26086957	-88.14910126	-85.31115723	2.83794403	15.4718357378
88478260.87	88.47826087	-87.27655029	-84.86638641	2.41016388	16.416706679
88695652.17	88.69565217	-86.06214142	-83.97612762	2.0860138	17.2199961125
88913043.48	88.91304348	-85.23097229	-83.33698273	1.89398956	17.7427109697
89130434.78	89.13043478	-84.77882385	-82.87493896	1.90388489	17.7147728264
89347826.09	89.34782609	-84.92500305	-82.85657501	2.06842804	17.2662548768
89565217.39	89.56521739	-85.13455963	-83.03114319	2.10341644	17.1745171429
89782608.7	89.7826087	-85.07232666	-82.68666077	2.38566589	16.4744467067
90000000	90	-85.47688293	-82.8739624	2.60292053	15.9769653057
90217391.3	90.2173913	-86.23091125	-83.91929626	2.31161499	16.6517343823
90434782.61	90.43478261	-87.01564026	-84.90028381	2.11535645	17.1434828006
90652173.91	90.65217391	-87.78918457	-85.85512543	1.93405914	17.6302876077
90869565.22	90.86956522	-88.47889709	-86.19418335	2.28471374	16.717202875
91086956.52	91.08695652	-88.66739655	-86.29697418	2.37042237	16.5105994247
91304347.83	91.30434783	-88.5428009	-86.84886932	1.69393158	18.3344241295
91521739.13	91.52173913	-88.52889252	-86.54077911	1.98811341	17.4815192574
91739130.43	91.73913043	-88.23061371	-86.27941895	1.95119476	17.5827744672
91956521.74	91.95652174	-87.94068146	-85.90507507	2.03560639	17.3534193657
92173913.04	92.17391304	-87.38462067	-85.67572021	1.70890046	18.2882432239
92391304.35	92.39130435	-87.54975128	-86.06497192	1.48477936	19.0177039062
92608695.65	92.60869565	-87.81601715	-86.22497559	1.59104156	18.661247554
92826086.96	92.82608696	-88.35749817	-86.56208801	1.79541016	18.027618602
93043478.26	93.04347826	-88.75837708	-86.92435455	1.83402253	17.9145619068
93260869.57	93.26086957	-88.68865204	-86.92099762	1.76765442	18.1101057846
93478260.87	93.47826087	-89.07777405	-86.98881531	2.08895874	17.2122792849
93695652.17	93.69565217	-88.74243927	-86.94528198	1.79715729	18.0224608846
93913043.48	93.91304348	-88.43480682	-86.46546936	1.96933746	17.532827794
94130434.78	94.13043478	-87.14611053	-85.8092804	1.33683013	19.5515335467
94347826.09	94.34782609	-85.08599091	-84.31391907	0.77207184	22.2295125755
94565217.39	94.56521739	-83.36557007	-83.24085999	0.12471008	30.4763232825
94782608.7	94.7826087	-82.73606873	-82.34602356	0.39004517	25.3902119815
95000000	95	-82.67516327	-82.53434753	0.14081574	29.9407329847
95217391.3	95.2173913	-83.83337402	-83.56602478	0.26734924	27.0926932315
95434782.61	95.43478261	-85.65799713	-84.79510498	0.86289215	21.6996905998
95652173.91	95.65217391	-87.24293518	-86.36238098	0.8805542	21.6025685955
95869565.22	95.86956522	-88.39296722	-86.92220306	1.47076416	19.066296988
96086956.52	96.08695652	-88.89265442	-87.56973267	1.32292175	19.6042627824
96304347.83	96.30434783	-89.03957367	-87.32330322	1.71627045	18.265627082
96521739.13	96.52173913	-88.9513092	-87.72137451	1.22993469	19.969551568
96739130.43	96.73913043	-89.05240631	-87.64598846	1.40641785	19.2945298018
96956521.74	96.95652174	-89.19297791	-87.42163849	1.77133942	18.0990942694
97173913.04	97.17391304	-89.12687683	-87.80116272	1.32571411	19.5936385352
97391304.35	97.39130435	-88.85306549	-87.69878387	1.15428162	20.2848091572
97608695.65	97.60869565	-89.0114212	-87.71372223	1.29769897	19.7011093439
97826086.96	97.82608696	-89.09892273	-87.58716583	1.5117569	18.9252407136
98043478.26	98.04347826	-89.04456329	-87.69592285	1.34864044	19.5071250576
98260869.57	98.26086957	-88.89644623	-87.66747284	1.22897339	19.9734505788
98478260.87	98.47826087	-88.76933289	-87.35739136	1.41194153	19.2745954376
98695652.17	98.69565217	-88.32752228	-87.1412735	1.18624878	20.1494686278
98913043.48	98.91304348	-87.94896698	-87.35267639	0.59629059	23.4417074918
99130434.78	99.13043478	-88.07015228	-87.13920593	0.93094635	21.3348117471
99347826.09	99.34782609	-88.14450836	-87.18043518	0.96407318	21.1657930534
99565217.39	99.56521739	-88.10877991	-87.23218536	0.87659455	21.6241883945
99782608.7	99.7826087	-88.58974457	-87.43765259	1.15209198	20.2941986408
100000000	100	-88.79283142	-87.54457855	1.24825287	19.8957523108

## ON
Machine Module	N9343C	CN0605A298
MCU Version	A.06.27
DSP Version	A.05.02.02 SA
FPGA Version	A.02.80
05:41:30 07.10.2020
Y Axis Scale	LOG
Y Axis Unit	dBm
Impedance	50
Number of Points	461
Sweep Time(s)	1.0051	MAN
Start Frequency	0
Stop Frequency	100000000
Frequency Offset	0
Average Type	Log Pwr
RBW	1000000	MAN
VBW	1000000	AUTO
Sweep Type	Continuous
Sweep Mode	Sweep
Speed	Fast
PreAmp State	OFF
Ref Level	-50
Attenuation	0	MAN
Ref Level Offset	0
X Axis Units	Hz
HiSensitive	OFF
Trigger Source	FreeRun
Trigger Level	0
Trigger Slope	RISE
Trigger Delay	0	OFF
Gate Sweep	OFF
Gate View	OFF
Gate Delay	0.00228
Gate Length	0.00084
Gate Source	External
Burst Level	-20
Burst Trigger Slope	Rise
Timer Period	0.00312
Timer Offset	0
Sync Source	Off
Apply Corrections	OFF
Antenna Unit	OFF
Correction1	Off
Correction2	Off
Correction3	Off
Correction4	Off
Marker Information
Marker	Trace	Type	X Axis	Amplitude
M1	1	Frequency	5.00 MHz	-65.35 dBm
M2	1	Frequency	10.00 MHz	-63.93 dBm
M3	1	Frequency	20.00 MHz	-64.43 dBm
M4	1	Frequency	30.00 MHz	-64.36 dBm
M5	1	Frequency	40.00 MHz	-64.62 dBm
M6	1	Frequency	50.00 MHz	-67.44 dBm
CISPR Filter	Normal RBW	Dwell Time(s)	0.5	AUTO
Trace Math	Off
Trace Math By	Log Power
Trace Variable A	Trace1
Trace Variable B	Trace2
Trace Variable C	Trace3
Trace Name	Trace1
Trace Count	20
Trace Type	Clear Write
Trace Detector	Average
Trace Data
0	-11.75434494
217391.3043	-27.07202148
434782.6087	-29.71936607
652173.913	-32.75317383
869565.2174	-36.8364563
1086956.522	-41.98287201
1304347.826	-48.27235413
1521739.13	-55.62693405
1739130.435	-62.23963547
1956521.739	-65.34380341
2173913.043	-65.85919189
2391304.348	-66.32976532
2608695.652	-66.56934357
2826086.957	-66.82597351
3043478.261	-66.65357971
3260869.565	-66.57684326
3478260.87	-66.22361755
3695652.174	-66.27230072
3913043.478	-65.95291138
4130434.783	-65.55252075
4347826.087	-65.44319153
4565217.391	-65.34928131
4782608.696	-65.21492004
5000000	-65.34703064
5217391.304	-65.05264282
5434782.609	-65.04060364
5652173.913	-65.13951111
5869565.217	-65.15223694
6086956.522	-64.95684814
6304347.826	-64.87095642
6521739.13	-64.69334412
6739130.435	-64.67814636
6956521.739	-64.50367737
7173913.043	-64.48991394
7391304.348	-64.40322113
7608695.652	-64.38274384
7826086.957	-64.41091919
8043478.261	-64.6634903
8260869.565	-64.33067322
8478260.87	-64.06374359
8695652.174	-64.29856873
8913043.478	-64.22408295
9130434.783	-64.34143829
9347826.087	-64.07247925
9565217.391	-63.77681732
9782608.696	-63.98671341
10000000	-63.92882538
10217391.3	-64.15724182
10434782.61	-63.77352142
10652173.91	-63.99614334
10869565.22	-64.01753235
11086956.52	-63.99403
11304347.83	-63.84386063
11521739.13	-64.33384705
11739130.43	-63.91174316
11956521.74	-64.2748642
12173913.04	-64.18322754
12391304.35	-64.15673065
12608695.65	-64.03100586
12826086.96	-63.99313354
13043478.26	-63.91299438
13260869.57	-64.03848267
13478260.87	-63.80094528
13695652.17	-63.78173065
13913043.48	-63.60549164
14130434.78	-63.75497437
14347826.09	-63.56461334
14565217.39	-63.82841873
14782608.7	-63.871521
15000000	-63.93060303
15217391.3	-63.99493408
15434782.61	-63.94477081
15652173.91	-63.96785736
15869565.22	-64.06208038
16086956.52	-64.16493225
16304347.83	-64.09477234
16521739.13	-64.23958588
16739130.43	-64.14215088
16956521.74	-64.05026245
17173913.04	-64.34202576
17391304.35	-64.47026825
17608695.65	-64.45251465
17826086.96	-64.11492157
18043478.26	-63.9429512
18260869.57	-64.25050354
18478260.87	-64.29623413
18695652.17	-63.94272614
18913043.48	-64.38813782
19130434.78	-64.49475098
19347826.09	-64.14842224
19565217.39	-64.28623962
19782608.7	-64.30835724
20000000	-64.43331909
20217391.3	-63.98868561
20434782.61	-64.25804138
20652173.91	-63.91373825
20869565.22	-64.16559601
21086956.52	-64.30021667
21304347.83	-64.20330048
21521739.13	-64.33314514
21739130.43	-64.20158386
21956521.74	-64.17578888
22173913.04	-64.28997803
22391304.35	-64.19422913
22608695.65	-64.30345917
22826086.96	-64.22190857
23043478.26	-64.01423645
23260869.57	-64.27472687
23478260.87	-64.27880859
23695652.17	-64.27531433
23913043.48	-64.33308411
24130434.78	-64.32089996
24347826.09	-64.32442474
24565217.39	-64.39015961
24782608.7	-64.07180023
25000000	-64.22285461
25217391.3	-64.14735413
25434782.61	-64.14304352
25652173.91	-64.27047729
25869565.22	-64.14464569
26086956.52	-64.2201004
26304347.83	-64.3344574
26521739.13	-64.1029129
26739130.43	-64.18247986
26956521.74	-64.21537781
27173913.04	-64.25626373
27391304.35	-64.12654114
27608695.65	-63.9965744
27826086.96	-64.0511322
28043478.26	-64.19047546
28260869.57	-64.0875473
28478260.87	-64.24711609
28695652.17	-64.21068573
28913043.48	-64.00536346
29130434.78	-64.23109436
29347826.09	-64.08595276
29565217.39	-64.02570343
29782608.7	-64.13220978
30000000	-64.35528564
30217391.3	-64.28220367
30434782.61	-64.27862549
30652173.91	-63.9835434
30869565.22	-63.98084641
31086956.52	-64.20365906
31304347.83	-64.05903625
31521739.13	-64.11010742
31739130.43	-63.99311066
31956521.74	-63.92526245
32173913.04	-63.90156937
32391304.35	-64.2824707
32608695.65	-64.289711
32826086.96	-64.25352478
33043478.26	-64.16029358
33260869.57	-64.02172089
33478260.87	-64.1255722
33695652.17	-64.19616699
33913043.48	-64.25237274
34130434.78	-64.05670166
34347826.09	-64.30054474
34565217.39	-64.1411972
34782608.7	-64.14465332
35000000	-64.21887207
35217391.3	-64.24980164
35434782.61	-64.01139069
35652173.91	-64.30021667
35869565.22	-64.3782959
36086956.52	-64.2589798
36304347.83	-64.29234314
36521739.13	-64.23854828
36739130.43	-64.26138306
36956521.74	-64.30053711
37173913.04	-64.50299072
37391304.35	-64.22431946
37608695.65	-64.47686005
37826086.96	-64.49191284
38043478.26	-64.28083801
38260869.57	-64.60081482
38478260.87	-64.30671692
38695652.17	-64.45336914
38913043.48	-64.48513794
39130434.78	-64.68629456
39347826.09	-64.69780731
39565217.39	-64.70810699
39782608.7	-64.66529846
40000000	-64.61640167
40217391.3	-64.79053497
40434782.61	-64.86800385
40652173.91	-64.89900208
40869565.22	-64.85600281
41086956.52	-64.71273804
41304347.83	-64.66358185
41521739.13	-64.79222107
41739130.43	-65.08765411
41956521.74	-64.71489716
42173913.04	-65.09364319
42391304.35	-65.00233459
42608695.65	-65.26438904
42826086.96	-65.2719574
43043478.26	-65.20870972
43260869.57	-65.08128357
43478260.87	-65.17927551
43695652.17	-65.3822403
43913043.48	-65.37236023
44130434.78	-65.29286957
44347826.09	-65.30142212
44565217.39	-65.29016113
44782608.7	-65.4553299
45000000	-65.62431335
45217391.3	-65.72602844
45434782.61	-65.75630188
45652173.91	-65.73703003
45869565.22	-65.93618011
46086956.52	-65.83976746
46304347.83	-65.89839172
46521739.13	-66.08417511
46739130.43	-66.13128662
46956521.74	-66.426651
47173913.04	-66.18965912
47391304.35	-66.49287415
47608695.65	-66.49877167
47826086.96	-66.53927612
48043478.26	-66.50498962
48260869.57	-66.65475464
48478260.87	-66.62548065
48695652.17	-66.76216125
48913043.48	-66.9750824
49130434.78	-67.04558563
49347826.09	-66.99937439
49565217.39	-67.27207947
49782608.7	-67.18508911
50000000	-67.44123077
50217391.3	-67.49909973
50434782.61	-67.48686981
50652173.91	-67.43919373
50869565.22	-67.61476135
51086956.52	-67.7460556
51304347.83	-67.72927856
51521739.13	-67.96530914
51739130.43	-67.93221283
51956521.74	-68.29960632
52173913.04	-68.22937012
52391304.35	-68.47247314
52608695.65	-68.52690125
52826086.96	-68.73858643
53043478.26	-68.8060379
53260869.57	-69.17069244
53478260.87	-69.16566467
53695652.17	-69.41433716
53913043.48	-69.51672363
54130434.78	-69.360672
54347826.09	-69.59672546
54565217.39	-69.81038666
54782608.7	-69.76555634
55000000	-69.70982361
55217391.3	-69.65032959
55434782.61	-69.76594543
55652173.91	-70.28158569
55869565.22	-70.53009796
56086956.52	-70.65022278
56304347.83	-70.83927155
56521739.13	-70.92736053
56739130.43	-70.98511505
56956521.74	-71.0994873
57173913.04	-71.39131165
57391304.35	-71.72305298
57608695.65	-71.74886322
57826086.96	-71.8001709
58043478.26	-71.93416595
58260869.57	-71.94932556
58478260.87	-72.35509491
58695652.17	-72.31001282
58913043.48	-72.27873993
59130434.78	-72.37924194
59347826.09	-72.44475555
59565217.39	-72.61538696
59782608.7	-72.74004364
60000000	-72.79734039
60217391.3	-72.7652359
60434782.61	-73.15041351
60652173.91	-73.41719818
60869565.22	-73.45783234
61086956.52	-73.65700531
61304347.83	-73.88663483
61521739.13	-74.03147888
61739130.43	-73.9972229
61956521.74	-74.33786774
62173913.04	-74.39678192
62391304.35	-74.47090912
62608695.65	-74.67082214
62826086.96	-74.96078491
63043478.26	-75.02131653
63260869.57	-74.92097473
63478260.87	-75.17274475
63695652.17	-75.24217224
63913043.48	-75.6192627
64130434.78	-75.75032806
64347826.09	-75.97353363
64565217.39	-76.13264465
64782608.7	-76.33377075
65000000	-76.58168793
65217391.3	-76.76039124
65434782.61	-76.70683289
65652173.91	-76.93788147
65869565.22	-77.22393799
66086956.52	-77.09860229
66304347.83	-77.01970673
66521739.13	-77.61719513
66739130.43	-77.59446716
66956521.74	-77.73770142
67173913.04	-77.74567413
67391304.35	-77.90193176
67608695.65	-77.96784973
67826086.96	-78.22543335
68043478.26	-78.24913025
68260869.57	-78.47307587
68478260.87	-78.5723114
68695652.17	-78.55595398
68913043.48	-78.93830872
69130434.78	-78.58622742
69347826.09	-78.93768311
69565217.39	-79.17084503
69782608.7	-79.07899475
70000000	-79.24319458
70217391.3	-79.54468536
70434782.61	-79.57946777
70652173.91	-79.64676666
70869565.22	-79.6175766
71086956.52	-79.98830414
71304347.83	-79.9885788
71521739.13	-80.00865936
71739130.43	-80.2374649
71956521.74	-80.26257324
72173913.04	-80.36961365
72391304.35	-80.32997131
72608695.65	-80.63053131
72826086.96	-80.76079559
73043478.26	-80.74305725
73260869.57	-80.89209747
73478260.87	-80.97000885
73695652.17	-81.21259308
73913043.48	-81.31587219
74130434.78	-81.2771759
74347826.09	-81.75141907
74565217.39	-81.3502655
74782608.7	-81.71173096
75000000	-81.5684967
75217391.3	-81.76560974
75434782.61	-81.75540161
75652173.91	-81.77290344
75869565.22	-82.02555084
76086956.52	-82.11621094
76304347.83	-82.2858963
76521739.13	-82.28256989
76739130.43	-82.46015167
76956521.74	-82.36006165
77173913.04	-82.35853577
77391304.35	-82.37187195
77608695.65	-82.59468842
77826086.96	-82.86076355
78043478.26	-82.8436203
78260869.57	-82.78552246
78478260.87	-82.75428772
78695652.17	-83.11578369
78913043.48	-83.14775848
79130434.78	-83.38368988
79347826.09	-83.07106781
79565217.39	-83.28617859
79782608.7	-83.64198303
80000000	-83.29314423
80217391.3	-83.30271912
80434782.61	-83.42510986
80652173.91	-83.8393631
80869565.22	-83.78521729
81086956.52	-83.99880981
81304347.83	-83.90058899
81521739.13	-83.80482483
81739130.43	-83.76229095
81956521.74	-84.13916016
82173913.04	-84.17308807
82391304.35	-84.0568924
82608695.65	-84.07229614
82826086.96	-84.30483246
83043478.26	-84.3225174
83260869.57	-84.66542053
83478260.87	-84.4048233
83695652.17	-84.29695892
83913043.48	-84.40657043
84130434.78	-84.60652924
84347826.09	-84.48709869
84565217.39	-84.77461243
84782608.7	-84.67166138
85000000	-84.86779022
85217391.3	-84.98560333
85434782.61	-84.9559021
85652173.91	-85.35970306
85869565.22	-85.04576111
86086956.52	-85.28408051
86304347.83	-85.23058319
86521739.13	-85.26670837
86739130.43	-85.51759338
86956521.74	-85.36219788
87173913.04	-85.52159882
87391304.35	-85.57054138
87608695.65	-85.63848877
87826086.96	-85.69380951
88043478.26	-85.50791931
88260869.57	-85.31115723
88478260.87	-84.86638641
88695652.17	-83.97612762
88913043.48	-83.33698273
89130434.78	-82.87493896
89347826.09	-82.85657501
89565217.39	-83.03114319
89782608.7	-82.68666077
90000000	-82.8739624
90217391.3	-83.91929626
90434782.61	-84.90028381
90652173.91	-85.85512543
90869565.22	-86.19418335
91086956.52	-86.29697418
91304347.83	-86.84886932
91521739.13	-86.54077911
91739130.43	-86.27941895
91956521.74	-85.90507507
92173913.04	-85.67572021
92391304.35	-86.06497192
92608695.65	-86.22497559
92826086.96	-86.56208801
93043478.26	-86.92435455
93260869.57	-86.92099762
93478260.87	-86.98881531
93695652.17	-86.94528198
93913043.48	-86.46546936
94130434.78	-85.8092804
94347826.09	-84.31391907
94565217.39	-83.24085999
94782608.7	-82.34602356
95000000	-82.53434753
95217391.3	-83.56602478
95434782.61	-84.79510498
95652173.91	-86.36238098
95869565.22	-86.92220306
96086956.52	-87.56973267
96304347.83	-87.32330322
96521739.13	-87.72137451
96739130.43	-87.64598846
96956521.74	-87.42163849
97173913.04	-87.80116272
97391304.35	-87.69878387
97608695.65	-87.71372223
97826086.96	-87.58716583
98043478.26	-87.69592285
98260869.57	-87.66747284
98478260.87	-87.35739136
98695652.17	-87.1412735
98913043.48	-87.35267639
99130434.78	-87.13920593
99347826.09	-87.18043518
99565217.39	-87.23218536
99782608.7	-87.43765259
100000000	-87.54457855

## OFF
Machine Module	N9343C	CN0605A298
MCU Version	A.06.27
DSP Version	A.05.02.02 SA
FPGA Version	A.02.80
05:42:48 07.10.2020
Y Axis Scale	LOG
Y Axis Unit	dBm
Impedance	50
Number of Points	461
Sweep Time(s)	1.0051	MAN
Start Frequency	0
Stop Frequency	100000000
Frequency Offset	0
Average Type	Log Pwr
RBW	1000000	MAN
VBW	1000000	AUTO
Sweep Type	Continuous
Sweep Mode	Sweep
Speed	Fast
PreAmp State	OFF
Ref Level	-50
Attenuation	0	MAN
Ref Level Offset	0
X Axis Units	Hz
HiSensitive	OFF
Trigger Source	FreeRun
Trigger Level	0
Trigger Slope	RISE
Trigger Delay	0	OFF
Gate Sweep	OFF
Gate View	OFF
Gate Delay	0.00228
Gate Length	0.00084
Gate Source	External
Burst Level	-20
Burst Trigger Slope	Rise
Timer Period	0.00312
Timer Offset	0
Sync Source	Off
Apply Corrections	OFF
Antenna Unit	OFF
Correction1	Off
Correction2	Off
Correction3	Off
Correction4	Off
Marker Information
Marker	Trace	Type	X Axis	Amplitude
M1	1	Frequency	5.00 MHz	-74.66 dBm
M2	1	Frequency	10.00 MHz	-72.72 dBm
M3	1	Frequency	20.00 MHz	-74.49 dBm
M4	1	Frequency	30.00 MHz	-74.64 dBm
M5	1	Frequency	40.00 MHz	-75.90 dBm
M6	1	Frequency	50.00 MHz	-78.47 dBm
CISPR Filter	Normal RBW	Dwell Time(s)	0.5	AUTO
Trace Math	Off
Trace Math By	Log Power
Trace Variable A	Trace1
Trace Variable B	Trace2
Trace Variable C	Trace3
Trace Name	Trace1
Trace Count	20
Trace Type	Clear Write
Trace Detector	Average
Trace Data
0	-11.99130917
217391.3043	-27.30935097
434782.6087	-29.95539856
652173.913	-32.98615265
869565.2174	-37.06858444
1086956.522	-42.22130585
1304347.826	-48.47438812
1521739.13	-55.87171555
1739130.435	-62.80449295
1956521.739	-66.6144104
2173913.043	-67.75583649
2391304.348	-69.43760681
2608695.652	-70.7430191
2826086.957	-70.80683899
3043478.261	-71.65588379
3260869.565	-72.16804504
3478260.87	-72.80612183
3695652.174	-73.09596252
3913043.478	-73.765625
4130434.783	-73.80515289
4347826.087	-74.30310059
4565217.391	-74.50592041
4782608.696	-74.59273529
5000000	-74.65511322
5217391.304	-74.70994568
5434782.609	-74.88475037
5652173.913	-74.99346161
5869565.217	-74.9924469
6086956.522	-74.87207031
6304347.826	-74.95466614
6521739.13	-74.82497406
6739130.435	-74.38682556
6956521.739	-74.11655426
7173913.043	-74.16059113
7391304.348	-73.55000305
7608695.652	-73.69324493
7826086.957	-73.96211243
8043478.261	-73.85462952
8260869.565	-73.31845093
8478260.87	-73.68213654
8695652.174	-73.45938873
8913043.478	-72.86204529
9130434.783	-72.8664093
9347826.087	-73.16951752
9565217.391	-72.32609558
9782608.696	-72.36161804
10000000	-72.72357941
10217391.3	-72.76444244
10434782.61	-72.85693359
10652173.91	-73.01094055
10869565.22	-72.47853088
11086956.52	-72.65557098
11304347.83	-72.9265976
11521739.13	-73.00440979
11739130.43	-73.00048828
11956521.74	-73.14477539
12173913.04	-73.51478577
12391304.35	-73.84142303
12608695.65	-73.99526978
12826086.96	-74.09633636
13043478.26	-74.36126709
13260869.57	-74.02961731
13478260.87	-74.24146271
13695652.17	-74.10066986
13913043.48	-74.26777649
14130434.78	-74.00718689
14347826.09	-73.93209839
14565217.39	-74.26480103
14782608.7	-74.39178467
15000000	-74.59461975
15217391.3	-74.71038818
15434782.61	-74.63832855
15652173.91	-74.81771088
15869565.22	-74.81835938
16086956.52	-75.04400635
16304347.83	-75.07546234
16521739.13	-75.33847046
16739130.43	-75.23096466
16956521.74	-75.4506073
17173913.04	-75.24313354
17391304.35	-75.53799438
17608695.65	-75.43787384
17826086.96	-75.32865143
18043478.26	-75.39024353
18260869.57	-75.46311951
18478260.87	-75.26559448
18695652.17	-75.30545044
18913043.48	-75.34015656
19130434.78	-75.08746338
19347826.09	-75.22006989
19565217.39	-75.03025818
19782608.7	-74.82463074
20000000	-74.48833466
20217391.3	-74.41654205
20434782.61	-74.49336243
20652173.91	-73.98191833
20869565.22	-73.74051666
21086956.52	-73.82462311
21304347.83	-73.49871063
21521739.13	-73.64199066
21739130.43	-73.95728302
21956521.74	-74.12381744
22173913.04	-74.09098816
22391304.35	-74.33403778
22608695.65	-74.46742249
22826086.96	-74.70526886
23043478.26	-74.82462311
23260869.57	-74.91353607
23478260.87	-74.80521393
23695652.17	-74.94657898
23913043.48	-75.04017639
24130434.78	-74.82086945
24347826.09	-74.75188446
24565217.39	-75.0558548
24782608.7	-74.99482727
25000000	-74.92945862
25217391.3	-74.98973083
25434782.61	-74.83776093
25652173.91	-74.78643799
25869565.22	-74.85609436
26086956.52	-75.16249084
26304347.83	-74.81100464
26521739.13	-74.73356628
26739130.43	-74.84461975
26956521.74	-74.98605347
27173913.04	-75.10969543
27391304.35	-74.99898529
27608695.65	-74.88171387
27826086.96	-74.7901535
28043478.26	-74.78247833
28260869.57	-75.02671814
28478260.87	-75.03601837
28695652.17	-74.9397583
28913043.48	-75.15668488
29130434.78	-74.87994385
29347826.09	-74.73290253
29565217.39	-74.80058289
29782608.7	-74.79386139
30000000	-74.64498138
30217391.3	-74.77867126
30434782.61	-74.67179871
30652173.91	-74.77639008
30869565.22	-74.89933777
31086956.52	-74.76811981
31304347.83	-74.87311554
31521739.13	-74.64003754
31739130.43	-74.93582916
31956521.74	-74.61646271
32173913.04	-74.94195557
32391304.35	-75.10417938
32608695.65	-74.88896942
32826086.96	-75.05574799
33043478.26	-74.76361847
33260869.57	-74.81978607
33478260.87	-74.95257568
33695652.17	-74.99253845
33913043.48	-74.99848938
34130434.78	-75.04120636
34347826.09	-74.85307312
34565217.39	-74.91560364
34782608.7	-74.96870422
35000000	-74.72984314
35217391.3	-74.93263245
35434782.61	-75.00539398
35652173.91	-75.22889709
35869565.22	-75.22747803
36086956.52	-75.31874847
36304347.83	-74.960495
36521739.13	-74.94078064
36739130.43	-75.1881485
36956521.74	-75.41871643
37173913.04	-75.26013947
37391304.35	-75.44788361
37608695.65	-75.53401184
37826086.96	-75.39095306
38043478.26	-75.50746155
38260869.57	-75.68080139
38478260.87	-75.31052399
38695652.17	-75.63084412
38913043.48	-75.48822784
39130434.78	-75.61436462
39347826.09	-75.5139389
39565217.39	-75.74758148
39782608.7	-75.91065979
40000000	-75.89892578
40217391.3	-75.5331192
40434782.61	-76.05297089
40652173.91	-75.79703522
40869565.22	-75.85343933
41086956.52	-76.15264893
41304347.83	-76.09310913
41521739.13	-75.92241669
41739130.43	-76.0559082
41956521.74	-76.38733673
42173913.04	-76.08969116
42391304.35	-76.03078461
42608695.65	-76.19051361
42826086.96	-76.12742615
43043478.26	-76.38376617
43260869.57	-76.3943634
43478260.87	-76.40690613
43695652.17	-76.55989838
43913043.48	-76.55091858
44130434.78	-76.84955597
44347826.09	-76.72579193
44565217.39	-76.90139771
44782608.7	-76.71944427
45000000	-76.96226501
45217391.3	-77.04482269
45434782.61	-76.99053192
45652173.91	-77.00837708
45869565.22	-77.12564087
46086956.52	-77.16051483
46304347.83	-77.29103088
46521739.13	-77.21916962
46739130.43	-77.16470337
46956521.74	-77.18651581
47173913.04	-77.39506531
47391304.35	-77.29965973
47608695.65	-77.6743927
47826086.96	-77.63111115
48043478.26	-77.66578674
48260869.57	-77.73119354
48478260.87	-77.9087677
48695652.17	-78.04768372
48913043.48	-78.27500153
49130434.78	-78.09738922
49347826.09	-78.24651337
49565217.39	-78.36375427
49782608.7	-78.37054443
50000000	-78.46876526
50217391.3	-78.53226471
50434782.61	-78.48509979
50652173.91	-78.25428772
50869565.22	-77.93986511
51086956.52	-77.88304901
51304347.83	-78.05119324
51521739.13	-78.25597382
51739130.43	-78.64247894
51956521.74	-78.80003357
52173913.04	-79.07688141
52391304.35	-79.35343933
52608695.65	-79.61909485
52826086.96	-79.40223694
53043478.26	-79.98161316
53260869.57	-79.75785828
53478260.87	-79.89911652
53695652.17	-80.10810089
53913043.48	-80.20404816
54130434.78	-80.21821594
54347826.09	-79.78277588
54565217.39	-78.76880646
54782608.7	-77.50254822
55000000	-76.25164032
55217391.3	-75.51876831
55434782.61	-75.90893555
55652173.91	-76.91920471
55869565.22	-78.35783386
56086956.52	-79.51385498
56304347.83	-80.75156403
56521739.13	-80.9524765
56739130.43	-81.31533051
56956521.74	-81.35207367
57173913.04	-81.44624329
57391304.35	-81.48139954
57608695.65	-81.73240662
57826086.96	-82.06419373
58043478.26	-82.22335052
58260869.57	-82.33327484
58478260.87	-82.20336151
58695652.17	-81.82099915
58913043.48	-81.16157532
59130434.78	-80.72842407
59347826.09	-80.55490875
59565217.39	-80.4171524
59782608.7	-80.53314972
60000000	-80.53755188
60217391.3	-80.38539886
60434782.61	-80.30888367
60652173.91	-80.29236603
60869565.22	-81.29347229
61086956.52	-81.4043045
61304347.83	-81.87042236
61521739.13	-82.00483704
61739130.43	-82.48990631
61956521.74	-82.66181946
62173913.04	-83.04972076
62391304.35	-83.32966614
62608695.65	-83.22769928
62826086.96	-83.47102356
63043478.26	-83.52866364
63260869.57	-83.23939514
63478260.87	-83.32798767
63695652.17	-83.76759338
63913043.48	-84.38884735
64130434.78	-84.53482819
64347826.09	-84.77827454
64565217.39	-85.33409119
64782608.7	-85.26881409
65000000	-85.46437073
65217391.3	-85.68286896
65434782.61	-85.75898743
65652173.91	-85.7155838
65869565.22	-86.09086609
66086956.52	-85.88470459
66304347.83	-86.05462646
66521739.13	-86.37687683
66739130.43	-86.10379028
66956521.74	-86.21642303
67173913.04	-86.25178528
67391304.35	-86.42797852
67608695.65	-86.29515076
67826086.96	-86.53401184
68043478.26	-86.79637909
68260869.57	-86.81448364
68478260.87	-86.78945923
68695652.17	-86.81301117
68913043.48	-86.81610107
69130434.78	-86.93493652
69347826.09	-87.29421234
69565217.39	-87.02078247
69782608.7	-87.40093231
70000000	-87.3057785
70217391.3	-87.40101624
70434782.61	-87.14692688
70652173.91	-87.52178955
70869565.22	-87.32141113
71086956.52	-87.73530579
71304347.83	-87.77687073
71521739.13	-87.80528259
71739130.43	-87.53786469
71956521.74	-87.83144379
72173913.04	-87.77987671
72391304.35	-87.79582214
72608695.65	-87.8584671
72826086.96	-87.76051331
73043478.26	-87.82221222
73260869.57	-87.79210663
73478260.87	-87.98830414
73695652.17	-87.79811096
73913043.48	-87.9997406
74130434.78	-87.80831909
74347826.09	-88.1288681
74565217.39	-88.09996033
74782608.7	-88.20426178
75000000	-88.0197525
75217391.3	-88.06745148
75434782.61	-88.02709198
75652173.91	-88.15932465
75869565.22	-88.36534119
76086956.52	-88.25410461
76304347.83	-88.22538757
76521739.13	-88.06748962
76739130.43	-87.91571808
76956521.74	-88.01524353
77173913.04	-88.03434753
77391304.35	-88.14778137
77608695.65	-87.93049622
77826086.96	-88.08961487
78043478.26	-88.01345825
78260869.57	-88.13181305
78478260.87	-87.87252045
78695652.17	-88.30793762
78913043.48	-88.07182312
79130434.78	-88.00430298
79347826.09	-88.29016113
79565217.39	-88.14643097
79782608.7	-88.03510284
80000000	-88.1505127
80217391.3	-88.13879395
80434782.61	-88.12400055
80652173.91	-88.19859314
80869565.22	-88.30750275
81086956.52	-88.04803467
81304347.83	-87.96338654
81521739.13	-88.15593719
81739130.43	-88.09835815
81956521.74	-88.21401215
82173913.04	-88.21116638
82391304.35	-87.99671936
82608695.65	-88.26070404
82826086.96	-88.30532837
83043478.26	-88.44459534
83260869.57	-88.25532532
83478260.87	-88.33422852
83695652.17	-88.28447723
83913043.48	-88.26490784
84130434.78	-88.4101181
84347826.09	-88.27347565
84565217.39	-88.35145569
84782608.7	-88.39717865
85000000	-88.42298126
85217391.3	-88.23323059
85434782.61	-88.50436401
85652173.91	-88.44972992
85869565.22	-88.56999207
86086956.52	-88.5273056
86304347.83	-88.59072113
86521739.13	-88.66267395
86739130.43	-88.5816803
86956521.74	-88.67098999
87173913.04	-88.61050415
87391304.35	-88.67906189
87608695.65	-88.5460434
87826086.96	-88.72386169
88043478.26	-88.48423004
88260869.57	-88.14910126
88478260.87	-87.27655029
88695652.17	-86.06214142
88913043.48	-85.23097229
89130434.78	-84.77882385
89347826.09	-84.92500305
89565217.39	-85.13455963
89782608.7	-85.07232666
90000000	-85.47688293
90217391.3	-86.23091125
90434782.61	-87.01564026
90652173.91	-87.78918457
90869565.22	-88.47889709
91086956.52	-88.66739655
91304347.83	-88.5428009
91521739.13	-88.52889252
91739130.43	-88.23061371
91956521.74	-87.94068146
92173913.04	-87.38462067
92391304.35	-87.54975128
92608695.65	-87.81601715
92826086.96	-88.35749817
93043478.26	-88.75837708
93260869.57	-88.68865204
93478260.87	-89.07777405
93695652.17	-88.74243927
93913043.48	-88.43480682
94130434.78	-87.14611053
94347826.09	-85.08599091
94565217.39	-83.36557007
94782608.7	-82.73606873
95000000	-82.67516327
95217391.3	-83.83337402
95434782.61	-85.65799713
95652173.91	-87.24293518
95869565.22	-88.39296722
96086956.52	-88.89265442
96304347.83	-89.03957367
96521739.13	-88.9513092
96739130.43	-89.05240631
96956521.74	-89.19297791
97173913.04	-89.12687683
97391304.35	-88.85306549
97608695.65	-89.0114212
97826086.96	-89.09892273
98043478.26	-89.04456329
98260869.57	-88.89644623
98478260.87	-88.76933289
98695652.17	-88.32752228
98913043.48	-87.94896698
99130434.78	-88.07015228
99347826.09	-88.14450836
99565217.39	-88.10877991
99782608.7	-88.58974457
100000000	-88.79283142